## Slide 1
ME486C FA25 – PIV Group
Santiago Gamez
Kian Pack
Christopher Rapoport
Jett Scruggs
Development of a Novel Particle Image Velocimetry System for the NAU Bioengineering Devices Laboratory
Research Area: Translational research on biomaterials, flow modeling, and medical device development for the treatment of aneurysms

## Slide 2: Design Requirement Summary
Chris 1

## Slide 3: Top Level Testing
Santiago 2
Experiment | Relevant DRs | Equipment Needed | Acceptance Criteria
Ex 1: Flow Study | CR2, CR3, CR4, CR5, CR7, ER2, ER3, ER4, ER5, ER7 and ER8 | Complete PIV Setup, pulsatile pump, tubing, particle solution, graduated cylinder, and stopwatch | System able to accurately analyze clinically relevant flow rates.
Ex 2: Safety Test | CR8 and ER9 | Complete PIV Setup | System meets ANSI Z.136
Ex 3: Laser Sheet Test | CR6 and ER6 | Laser assembly, blank wall, a ruler, protective goggles | Laser Sheet Width >1 mm
Ex 4: Timed Setup Test | CR1 and ER1 | Complete PIV Setup, timer | Setup Time < 60min

## Slide 4: Set-up Test – Procedure & Results
Jett 3
Setup Time Success Metric:
Set up and adjust the PIV setup in under 60 minutes
Fulfills requirement for ease of use
Procedure:
PIV setup is initially positioned as if it were stored
Timer is started, and the components are positioned around the flow model being studied
Once positioned, the components are adjusted to the model
Final adjustments are made, and the timer is stopped
Results:
Combined positioning and adjusting took 20 minutes total
Time of setup is far less than maximum time limit, exceeds goal

## Slide 5: Safety Test - Procedure
Kian 4
Class 3B Laser Definition: "A class 3B laser may be hazardous under direct and specular reflection viewing conditions, but is not normally a diffuse reflection or fire hazard."
Evaluated Safety Features:
Tool Access Master Arm Switch
Positively Actuated Shutter
Accessible Interlocked Emergency Stop
Multiple Shut-Down Controls
Visual Indication of Operation
Extant Safety Controls:
Laser Room Isolation
Proper PPE
Proper Signage

## Slide 6: Safety Test – PPE Evaluation
Kian 5
Important Laser Characteristics:
Calculated Required OD: 0.68
LG14C OD at 520nm: >2.0

## Slide 7: Safety Test - Videos
Kian 6

## Slide 8: Flow Study
Chris 7
The flow study will determine whether the system is capable of analyzing flow data at the velocities seen in the BDL.
  Camera must be able to capture images of adequate quality at a high enough frame rate analysis
  The laser must be able to illuminate the particles enough to be tracked by PIVlab
  The function generator must be able to synchronize the laser and camera
A flow loop consisting of a pulsatile pump, ½" OD tubing and a reservoir of PIV solution will be analyzed with our PIV system.
The flow system will be set to a flow rate of 240 mL/min to match the velocities in the BDL flow model.
3 Bucket timers will be taken to confirm the flow rate.
Images of the flow will be taken at 1,000 Hz, with the laser being pulsed before and after each frame.
Images will be uploaded to PIVlab for analysis.
DRs Tested: CR2, CR3, CR4, CR5, CR7, ER2, ER3, ER4, ER5, ER7 and ER8

## Slide 9: Light Sheet test – Procedure
Santiago 8
Light Sheet Success Metric:
Light sheet width of more than 1 mm at operating distance
Procedure:
All standard safety precautions are used for setting up PIV system- only two operators in closed room
Laser is aimed at non-reflective surface
High precision measuring instruments are positioned next to laser beam profile
High resolution camera with sufficiently low exposure setting captures image of beam profile
Measure beam profile width from images
Further Actions​:
If width is out of spec, position of cylindrical lens will be adjusted until in desired range

## Slide 10: CR Table
Santiago 9
Customer Requirement | CR Met Y/N | Client Acceptable Y/N
CR1: Easy to set up and position camera and laser | Y | Y
CR2: High enough camera resolution | Y | Y
CR3: High enough camera framerate | Y | Y
CR4: Camera runtime long enough to capture phenomena of interest | TBD | TBD
CR5: High enough laser power | TBD | TBD
CR6: Light sheet wide enough | TBD | TBD
CR7: Function generator fast and accurate enough | Y | Y
CR8: System safety | Y | Y

## Slide 11: ER Table
Kian 10
Engineering Requirements | Target | Tolerance | Measured or Calculated Value | ER Met Y/N | Client Acceptable Y/N
ER 1: Setup Time | ≤ 60 min | N/A | 20 min | Y | Y
ER2: Camera Resolution | ≥ 1 MP | N/A | 0.262 MP | N | Y
ER3: Camera Framerate | ≥ 200 FPS | N/A | 5,000 FPS | Y | Y
ER4: Camera Runtime | ≥ 10 s @ 200 FPS | N/A | TBD | TBD | TBD
ER5: Laser Energy | ≥ 100 nJ | N/A | TBD | TBD | TBD
ER6: Light Sheet Width | ≥ 1 mm | N/A | TBD | TBD | TBD
ER 7: Function Generator Resolution | ≤ 500 ps | N/A | 250 ps | Y | Y
ER8: Function Generator Trigger Rate | ≥ 20 kHz | N/A | 1 GHz (manufacturer spec.) | Y | Y
ER9: System Safety | ANSI z136 | N/A | In Compliance with ANSI Z.136 | Y | Y

## Slide 12: QFD
Jett 11

## Slide 13: Next Testing Steps
Chris 11
Flow Study: Ensure that the PIV system can capture images at high enough speeds and of sufficient quality for high-fidelity PIV analysis
Light Sheet Test: Ensure that the width of the laser light sheet meets our client's requirements

## Slide 14: Thank you! Questions?