## Slide 1
ME476C SP25 – PIV Group
Santiago Gamez
Kian Pack
Christopher Rapoport
Jett Scruggs
Development of a Novel Particle Image Velocimetry System for the NAU Bioengineering Devices Laboratory
Research Area: Translational research on biomaterials, flow modeling, and medical device development for the treatment of aneurysms

## Slide 2: Introduction
Particle Image Velocimetry (PIV):
A system that uses a combination of a laser and an ultra-high-speed camera to track the change in position of fluorescent tracer particles within a fluid flow, allowing the analysis of the velocity field within the fluid.
Chris 1
Fig. 1. Overview of Typical PIV System [1]
Fig. 2. Previous PIV study conducted by the BDL

## Slide 3: Project Overview
Sponsor: Tim Becker, PhD
Associate Professor
Mechanical Engineering Department
Principle Investigator
Bioengineering Devices Laboratory (BDL)
Chief Technology Officer
Aneuvas Technologies, Inc. (ATI)
Chris 2
Client: Wyatt Clark, M.S.
PhD Candidate
Mechanical Engineering Department
Senior Research Associate
Bioengineering Devices Laboratory (BDL)

## Slide 4: Project Justification
Almost 800,000 strokes occur in the US every year [2].
Ischemic Stroke: blood flow is blocked in the brain – treated with drugs or mechanical thrombectomy.
Hemorrhagic Stroke: rupture of blood vessel in the brain – treated with coiling and stents.
Endovascular treatments currently about 60% effective [3].
Chris 3
Fig. 3. BDL aspiration thrombectomy study (top) BDL coiling study (bottom)

## Slide 5: Design Description : Lens & Laser Configuration
Laser turns on
Convex lens contracts beam
Concave Lens expands beam
Cylindrical lens to diverge beam in x-dimension
Cylindrical lens to diverge beam in y-direction
Santiago 4

## Slide 6: Design Description : Convex Lens Mount
Santiago 5
Mount holds convex lens in place

## Slide 7: Design Description : Concave Lens Mount
Mount holds concave lens in place
Santiago 6

## Slide 8: Design Description : Breadboard
Holds all mounting systems in one uniform line
Santiago 7

## Slide 9: Design Description : Laser
Santiago 8
Start: Laser passes through lenses to create sheet size of 0.1 mm
End: Lights up fluorescent particles in flow

## Slide 10: Customer and Engineering Requirements
Customer Requirements
Ease of Setup
Camera Resolution
Camera Frame Rate
Camera Run Time
Strength of Laser
Adherence to Laser Safety Standards
Acceptable Width of Laser Sheet
Use of Multifunction Generator
Jett 9
Engineering Requirements
Setup Time (<60min)
Resolution (~1.3MP) (From ~2.1MP)
Camera Frame Rate (200FPS) (From 1000FPS)
Run Time (10s)
Laser Power (50-150mW)
Risk Assessment Score (1-10)
Sheet Thickness (100-200µm)
Frequency (10ns)
Refractivity Index (1.5RI)

## Slide 11: Updated QFD
Jett 10
Table 1: QFD

## Slide 12: Engineering Analysis – Optimization of Camera and Lens Setup
Jett 11
Optimization helps to detail how well certain specifications will work together, and what to change
Camera lens has an assumed magnification of 15, a sensor length of 11.3mm.
Laser wavelength of 520nm, a numerical aperture of 2 and medium refractive index of 2.8.
Reference of flow was the Basilar artery, with an average flow velocity of 0.55m/s [5]

## Slide 13: Engineering Analysis – F of Lenses & Length
Santiago 12
(1)
(2)
(3)
Figure 4: Two lenses in configuration [6]

## Slide 14: Engineering Analysis – Aneurysm Model
Kian 13
Boundary Conditions:
  Velocity Inlet
    Specified at 0.23 m/s based on empirical data [6]
  Pressure Outlet
    Based on averaged healthy systolic and diastolic pressures – 100mmHg from 120/80 mmHg.
      Set to 13,332 Pa in CFD
  Wall Conditions
    No-slip condition as deformation and wall shear stress are ignored in this case.
Material Properties:
  Fluid treated as water for this simulation.
  Particles treated as a massless discrete phase.
0.23
Table 2. Boundary Conditions for CFD as set in ANSYS Fluent.

## Slide 15: Engineering Analysis – Optical Density for Safety Goggles
Chris 15
Use ANSI Z.136.1 Tables 5-7 to find equations based on laser specifications and application
Required OD – 4.69

## Slide 16: Design Validation - FMEA
Laser Potential Modes of Failure:
Diode Overheating
  Engineering Controls – Integrated temperature controller in driver.
Diode Overcurrent
  Engineering Controls – Overcurrent protection circuit in driver.
Electro-Static Discharge (ESF) to Electronic Components
  Elimination via Process Control – Operator must ground before interacting with laser.
  PPE – Operators must wear anti-ESD wristbands when interacting with laser
Camera:
Overheating
  Engineering Control – Temperature sensor in camera
Overcurrent
  Engineering Control – Overcurrent protection circuit in camera
Kian 16

## Slide 17: Design Validation - Testing
Performing Experiment with Laser
Timed setup
Comparison of flow studies to CFD and literature
Testing laser safety measures
Testing equipment: BDL comprehensive benchtop model
Chris 17
Fig. 5. Top: Laser Fluorescence Experiment Bottom: BDL Comprehensive Flow Model

## Slide 18: Project Schedule
Kian 18

## Slide 19: Budget and Plan
Kian 19
Spend for SP25
Critical Equipment-
ThorLabs NPL52C Pulse Laser -$2,300
R&D Spend (Capstone Budget)-
Laser Safety Goggles: -$538
CLD1010LP DC Power Supply: -$85
Budget Buckets Remaining
Critical Equipment: $12,700
R&D: $2,377
Fundraising Updates
BDL Equipment Donation-
ThorLabs CLD1010LP laser driver and diode combo: ~$3,000
Hooper Undergraduate Research Grant-
Total Potential Equipment Funding: $1,500
Remaining Equipment
Camera (Filter+Lenses)
Mounting Equipment
Laser Optics Train
Function Generator

## Slide 20: References
20
[1] Seika Digital Image Corporation, “PIV ( Particle Image Velocimetry ) | Overview & Principle.” [Online]. Available: https://www.seika-di.com/en/measurement/principle_of_piv.html
[2] M. Katan and A. Luft, “Global Burden of Stroke,” Semin Neurol, vol. 38, no. 02, pp. 208–211, Apr. 2018, doi: 10.1055/s-0038-1649503.
[3] “How many people are affected by/at risk for stroke? | NICHD - Eunice Kennedy Shriver National Institute of Child Health and Human Development.” Accessed: Feb. 26, 2025. [Online]. Available: https://www.nichd.nih.gov/health/topics/stroke/conditioninfo/risk
[4] F. Zigunov, “How bright are my piv particles?,” Zigunov Aero, https://zigunov.com/2023/06/30/how-bright-are-my-piv-particles/ (accessed Mar. 3, 2025).
[5] Allen D Elster, “Expected Velocities,” Questions and Answers in MRI. [Online]. Available: https://mriquestions.com/expected-velocities.html
[6] “Light sheet optics for PIV - HackMD,” HackMD, 2023. https://hackmd.io/@cdcl-rpr/light-sheet-optics

## Slide 21: Virtual Prototype Results
13
Fig. 6. Vector Plot of Simulated Test Flow

## Slide 22: Virtual Prototype Results - Cont'd
13
Fig. 7. Pathlines Graphic of Simulated Test Flow

## Slide 23: Thank you! Questions?